Pacemakers and AICD

- how they work and why they sometimes don’t

Pacemaker nomenclature

I	- chamber paced  (Ventricle – Atria – Dual)

II	- chamber sensed  (V – A – D)

III	- response to sensing  (Trigger, Inhibit, Dual)

For example, common pacemakers include:

VVI – paces ventricle, sensing ventricular beat causes inhibition of pacing

DDD – senses and paces both chambers (2 leads), physiologic

VVI pacemaker

VVI pacemaker

DDI pacemaker

Unipolar lead – electricity between tip of lead and generator  large voltage spike

Bipolar lead – electricity between two electrodes on tip of lead  small voltage spike

First quiz – why is this pacemaker firing at 130/minute?

Pacemaker malfunction

Failure to pace or discharge impulse

- component failure – rare

- total battery depletion – VVI 10-15 years, DDD 7-10 years

- lead fracture or disconnection

- oversensing – muscular activity or external electrical interference

Failure to capture myocardium

- lead dislodgement – most common cause

- lead insulation break

- elevated myocardial threshold

- battery depletion

Failure to sense spontaneous depolarizations

- lead dislodgement

- poor lead placement

- change in cardiac signal or QRS morphology

2nd Quiz – Elderly woman with syncope in church

Quiz #3 – Elderly man with palpitations

Specifics of pacemaker malfunction

Lead dislodgement

- usually within first months after implantation

- increased risk – large RV, poor placement

- “Twiddler’s syndrome” – unintentional manipulation of generator

- into SVC or right atrium  phrenic nerve stimulation  pacing of diaphragm

hiccups – change in pacer spike and QRS complex

Lead fracture

- usually months after implantation

- caused by stress at connection site, chest trauma

Lead perforation

- usually soon after implantation (80% within 4 days)

- may cause acute tamponade

- failure to capture and sense, or change in complex (LBBB  RBBB)

Failure to capture 2° elevated myocardial threshold

- threshold gradually increases 2° fibrosis

- hyperkalemia, acidosis, hypoxemia

- ischemia or infarction

- antiarrhythmic drug toxicity

Hysteresis

- programmable function that allows rate to be slightly lower than set

- may allow patients to retain their own rhythm without battery depletion

- EKG – inappropriately long interval after intrinsic beat

1

2

3

4

5

6

7

8

9

10

11

12

ICD – Implantable cardioverter-defibrillator

Pacemaker with defibrillation coils

Complications

- inappropriate detection and delivery of shock – most common

- wound/pocket problems  (hematoma, infection)

- lead dislodgement or migration

Inappropriate shocks (in up to 40% patients)

- interrogate the device - reprgramming

- usually SVT (including sinus tach, atrial fib), also electrical noise, lead fracture

- better results with modern programming strategies

Phantom shocks – (often have received prior real shock causing anxiety)

Appropriate shocks

- rate of appropriate shocks 2-3X rate of sudden death

- maybe shocked some rhythms that would not have been sustained

- add medication if frequent appropriate shocks

Electromagnetic interference – surveillance systems, MRI

Antitachycardia pacing – first attempt, should terminate 90% v tach

EMS concerns – CPR, external defibrillation

13